## Slide 1
Effects of Spacer Devices in pMDI Drug Delivery with Ansys Fluent®
Developed and curated by the Ansys Academic Development Team
Juan Ignacio Doval
education@ansys.com
Inspired by Kleinstreuer Spacer Model [1] and Alberta Idealized Throat [2]

## Slide 2: Learning Objectives
Intended Learning Outcomes | 
Knowledge and Understanding | Understand the primary Physics of Discrete Phase Modelling and Species Transport for Drug Delivery (pMDI)
Skills and Abilities | Know how set up a Multiphysics Simulation for Continuous Phase and Discrete Phase Modeling
Values and Attitudes | Appreciate the importance of In-Silico technologies for Drug Delivery Studies Develop critical thinking when interpreting CFD results in a biomedical context
Resources
Built-in Assessment
Further reading/information
  Ansys Academic
  Ansys Innovation Courses
  Ansys Education Resources
Ansys software mentioned | Ansys Fluent® fluid simulation software Ansys Discovery™ 3D product simulation software

## Slide 3
Pressurized metered-dose inhalers (pMDIs) are widely used for respiratory drug delivery but often suffer from poor hand–breath coordination, which can lead to high oropharyngeal deposition and reduced lung delivery efficiency. Spacer devices help mitigate this issue by providing an expansion chamber where the aerosol plume decelerates and the propellant partially evaporates due to the rapid pressure drop. This process reduces particle velocity and size, allowing finer particles to be inhaled more deeply into the lungs and minimizing oropharyngeal losses. This study presents a computational proof of concept demonstrating the use of Ansys software for Drug Delivery Simulations.
/ About
pMDI + Idealized Throat Model
pMDI + Spacer + Idealized Throat Model

## Slide 4
Step-by-Step Instructions for Multi-physics Ansys Fluent Set-up
Content: Demonstration of effects of the Spacer Chamber in Drug Delivery using Pressurized metered-dose inhalers (pMDI)
/ How To Use This Tutorial

## Slide 5
Section 1 – Multiphysics Simulation Set-Up
Solver:
Transient Species Transport & DPM Model
Post-Processing for DPM
/ Software Overview

## Slide 6
/ Respiratory Drug Delivery Challenge
Problem Statement:
300+ million people worldwide suffer from asthma
Poor coordination during pMDI use → high oropharyngeal deposition
Only about 8-53% of drug reaches lungs without spacer due to poor hand-lung coordination, with worse results in asthmatic patients (https://pmc.ncbi.nlm.nih.gov/articles/PMC3123836/)
A common Solution:
Spacer devices as add-on accessories

## Slide 7
/ What is a pMDI?
Pressurized Metered-Dose Inhalers consist of:
Components:
Canister
Metering valve
Actuator nozzle
Operation:
High-pressure release (10 atm) → Plume Formation
Spray Characteristics:
Exit velocity: 80-225 m/s
Initial droplet size: >20 μm
Final particle size: ~1 μm (post propellent evaporation)
Canister
Drug/Propellant
Metering Valve
Actuator
Nozzle
Spray/Plume
CAD Model

## Slide 8
/ Evolution of pMDI Technology
The timeline of pMDI technology includes:
1956: First pMDI developed
1990s: CFC phase-out → HFA-134a propellants
2000s: Smaller nozzles (0.5mm → 0.25mm)
Present: Spacer optimization & smart delivery systems
Key Drivers:
Improved drug delivery
Environmental regulations

## Slide 9
/ Relevant Analytical Models
The simulation employs the following fundamental models to accurately predict particle behaviour:
Particle Dynamics
  Stokes Drag Law
  Particle Relaxation Time
  Gravitational Settling
  Turbulent Dispersion
  Species Transport
Deposition Mechanism
  Inertial Impaction
  Stokes Number
  Turbulent Deposition
Geometry Influence
Press the bottom to take part in a guided learning course about Fluid Dynamics in Ansys Fluent

## Slide 10: Particle Dynamics & Flow Physics Theory

## Slide 11
/ Particle Dynamics
Particle motion is governed by three primary force categories acting on individual droplets as they traverse the computational domain:
Stokes drag provides the dominant resistance force, with drag coefficient varying based on particle Reynolds number to account for non-linear effects at high relative velocities. The characteristic particle response time evolves as particles decelerate from initial injection velocities to air flow velocities, transitioning from intermediate drag regime to Stokes regime.
Gravitational settling contributes to vertical particle motion, with terminal velocity proportional to the square of particle diameter, becoming particularly important in low-velocity regions where residence times are extended.
Turbulent dispersion is modeled using the Discrete Random Walk (DRW) approach, where particles interact with random turbulent eddies defined by the time scale constant, providing stochastic velocity fluctuations that scatter particles toward walls and enhance deposition probability.

## Slide 12
/ Stokes Drag Law
Particle-Fluid Momentum Exchange
Relevance
Primary force governing particle deceleration in air flow
Determines particle trajectory curvature and response to flow changes
Critical for predicting particle behavior in spacer geometry
or
Viscosity of Fluid
radius of Particle
particle velocity relative to fluid
Drag Force
Flow-Induced Force
Gravitational Force
Description
Stokes drag law describes the viscous resistance force acting on spherical particles moving through a fluid at low Reynolds numbers. While it is often discussed in the context of gravitational settling, the same principle applies to any motion through a viscous medium. In the case of aerosolized particles from a pMDI entering a spacer, Stokes drag represents the viscous resistance that slows particles as they decelerate from high injection velocities within the chamber.

## Slide 13
/ Particle Relaxation Time
Particle-Fluid Momentum Exchange
Description
The particle relaxation time represents the characteristic time scale for a particle to adjust its velocity in response to changes in the surrounding flow field. It indicates how quickly particles can follow or react to flow changes.
Relevance
Determines particle inertia and response to flow acceleration
Key parameter for predicting impaction at geometric features
Foundation for calculating Stokes number
Critical for understanding particle “lag” behind air flow
Higher
Lower
i.e., reactive to fluid flow changes
i.e., unreactive to fluid flow changes
Salt
Dust
Smoke

## Slide 14
Movement of Particles: Newton’s Second Law:
If Drag is considered:
If flow around the particle is in Stokes Regime:
Where: – Fluid Velocity
& – Particle Velocity
i.e., Driven by the discrepancies of velocities between particle and fluid flow direction.
Then:
Substituting & Balancing:
Particle Response Time (Stoke Flow)
Where:
Only for
/ Particle Relaxation Time
Particle-Fluid Momentum Exchange

## Slide 15
For this case:
@ Spray Nozzle m/s & , therefore
So, for
And,
Therefore, quadratic in velocity difference , This is a Non-Linear ODE
Where particle relaxation time is dependent on which itself depends on .
Stokes: :
Intermediate: :
It needs to be solved Numerically!
/ Particle Relaxation Time
Particle-Fluid Momentum Exchange
NOTE
How fast the particle accelerates or decelerates at a given time?
How long it takes for the particle to travel at air speed?

## Slide 16
/ Gravitational Settling (Terminal Velocity)
Body Forces
Description
Gravitational settling describes the steady-state velocity reached by particles when gravitational force balances drag and buoyancy forces. This represents the particle's natural falling velocity in still air.
Relevance
Important for vertical deposition in spacer geometry
Affects residence time of particles in spacer
Determines settling patterns in low-velocity regions
Critical for longer spacers where settling becomes significant

## Slide 17
Substituting & Balancing equations
General Terminal Velocity Equation
/ Gravitational Settling (Terminal Velocity)
Body Forces
At Stokes Regime:
And Terminal Velocity becomes:

## Slide 18
This features is added into the model by enabling gravity, by providing the Gravity vector such as [0, -9.81, 0], since Y is the vertical axis.
As shown in the figure: Task Page > Gravity >
/ Gravitational Settling (Terminal Velocity)
Body Forces
NOTE: The component is dependent to the orientation of your Discovery CAD file, for this case –Y is vertically down.

## Slide 19
/ Turbulent Dispersion (DRW)
Stochastic Particle Dispersion
Description
Turbulent dispersion models the random scattering of particles by turbulent eddies through the Discrete Random Walk (DRW) approach. Particles interact with successive random eddies, each defined by a lifetime, length scale, and velocity magnitude, providing stochastic velocity fluctuations that enhance wall deposition probability.
Relevance
Provides random "kicks" that scatter particles toward walls
Essential for flow-following particles to reach deposition surfaces
Time scale constant controls mixing intensity and deposition rates
Captures collective turbulent mixing effects without explicit collision modelling
Eddy Lifetime:
Random velocity:
Particle Dispersion:
More can be found here!

## Slide 20
Description
The Discrete Phase Model is a computational approach capable of tracking the motion of individual particles—such as drug aerosols, droplets, or solid particles within a continuous fluid flow. In DPM, particles are treated as distinct entities that interact with the fluid through various forces including drag, lift, gravitational, and turbulent dispersion forces, as well as wall interactions. This approach provides detailed information about particle distribution, velocity fields, and individual trajectories throughout the flow domain.
Mathematically, DPM employs Langrangian particle tracking, a framework that follows individual particles through the flow field by solving Newton's second law along particle trajectories. The governing equations are:
Position Equation
Momentum Equation (Newton’s 2nd Law)
Expanded Form
Stokes Drag
Gravity
/ Discrete Phase Model (DPM)
Langrangian Particle Tracking

## Slide 21
/ Species Transport
Multi-Component Mass Transfer
General Form of Species Transport Equation:
Description
Species transport governs the continuous-phase mass transfer of mixture components by solving equations for convection and diffusion. When coupled with DPM, a source term () is added to represent propellant vapor released from evaporating droplets, which is then tracked as it disperses through the spacer.
Component 2
API
Component 1
Propellant
t = 0
t
Evaporation of Component 1
Where:
Rate of Change of density weighted mass fraction for species .
Advection term, representing the transport of the species by the fluid flow.
Diffusion term, where is the diffusion coefficient of the species, modelling its transport due to molecular motion
Mass added to gas phase from all evaporating particles

## Slide 22: Deposition Mechanism Theory

## Slide 23
Description
The Stokes number is a dimensionless parameter that characterizes the ratio of particle inertial forces to viscous forces. It indicates whether particles will follow the air flow or maintain their original trajectory due to inertia.
Fundamental Form:
Response Time x Flow Rate
Stokes Number Regimes:
: Flow-following (Viscous-like)
: Transitional behavior
: Ballistic (Inertial-like)
Where: – Fluid Velocity
& – Characteristic Length Scale
Explicit Form:
/ Stokes Number
Inertial Impaction

## Slide 24
/ Inertial Impaction
Ballistic Deposition Mechanism
Description
Inertial impaction occurs when particles with high inertia cannot follow sudden flow direction changes and collide with walls. This mechanism is characterized by the Stokes number, which represents the ratio of particle response time to flow time scale.
Relevance
Dominant deposition mechanism for large particles (>5 μm)
Most effective at geometric discontinuities and sharp bends
Primary cause of oropharyngeal deposition in pMDI applications
Efficiency increases rapidly with Stokes number
Stokes Number (St):
At Stokes Regime ()
Impaction Efficiency (n):
For Cylindrical Geometries
Where: – Fluid Density, – Particle Diameter
– Characteristic Length Scale
– Characteristic Velocity
& – Dynamic Viscosity
Particle ID: 7
Particle Diameter: 16 µm
Deposited due to Ballistic Impaction

## Slide 25
/ Turbulent Deposition
Transport-Limited Deposition
Description
Turbulent deposition occurs when turbulent eddies transport particles from the bulk flow toward walls through enhanced diffusion and mixing. Unlike laminar conditions where particles follow smooth streamlines, turbulence creates cross-stream velocity fluctuations that increase particle-wall contact probability, particularly important for smaller particles that would otherwise remain suspended.
Particle ID: 17
Particle Diameter: 16 µm
Deposited due to Turbulent Dispersion
High Turbulent Kinetic Energy () Region

## Slide 26: Geometry & Boundary Effects Theory

## Slide 27
/ Flow Separation & Patterns
Flow Field Characteristics
Description
Flow separation occurs at sharp geometric transitions, creating recirculation zones and regions of flow reattachment. These flow patterns significantly influence particle transport by creating low-velocity regions where settling becomes important and high-shear zones that enhance deposition.
Relevance
Creates Particle Trapping zones downstream of constrictions
Modifies local Stokes numbers through velocity variations
Establishes preferential deposition locations
Recirculation Region
Separation point
Preferential Particle Deposition Location Downstream
Reattachment Region
Constriction – Small Cross-Section Area

## Slide 28
/ Throat Constriction
Geometric Design Effects
Description
Throat constriction geometry determines local flow acceleration, pressure gradients, and particle trajectory deflection. The constriction ratio, entrance angle, and downstream expansion significantly influence deposition patterns through their effect on local Stokes numbers and flow separation characteristics.
Constriction ratio:
Where
Reynolds number:
Pressure drop:
Stokes Number:
Particle Deposition @

## Slide 29: pMDI + Idealized Throat Model

## Slide 30
/ Lung Deposition Pt.1
pMDI + Idealized Throat
This study examines two computational models: pMDI + Idealized Throat and pMDI + Spacer + Idealized Throat. The first configuration serves as a validation case for the multi-physics implementation, provides baseline deposition characteristics and serves as a demonstrational model to study the physics involved.
Regarding lung delivery efficiency, experimental literature indicates that for pMDI + Idealized Throat configurations, only 8-53% of the active pharmaceutical ingredient (API) dose successfully reaches the lungs, with the remainder depositing in the upper respiratory tract. To capture realistic patient breathing patterns, three clinically relevant inhalation flow rates were simulated: 30 L/min (light breathing), 60 L/min (moderate breathing), and 90 L/min (strong breathing), representing the spectrum of patient inspiratory efforts typically observed in clinical practice.
NOTE: It should be noted that for this tutorial, particles that escape through the computational domain outlet are considered successful lung deposition, as the outlet boundary represents the tracheal terminus where particles would continue toward the lower respiratory tract.

## Slide 31
/ Lung Deposition Pt.2
pMDI + Idealized Throat – Results
The results collected from the three inhalation flow rates (30 L/min, 60 L/min & 90 L/min) are the following:
Results & Validation |  |  |  |  | 
Flow Rates () | Injected Particles | Trapped Particles | Escaped Particles | Model Efficiency | Experimental Efficiency
30 | 1000 | 496 | 504 | 50.4% | 8-53%
60 | 1000 | 731 | 269 | 26.9% | 
90 | 1000 | 800 | 200 | 20.0% | 
Results & Validation |  |  |  |  | 
 | Injected Particles | Trapped Particles | Escaped Particles | Model Efficiency | Experimental Efficiency
30 | 1000 | 496 | 504 | 50.4% | 8-53%
60 | 1000 | 731 | 269 | 26.9% | 
90 | 1000 | 800 | 200 | 20.0% | 
NOTE: In literature, 2 more flow rates are occasionally considered: 15 L/min & 120 L/min.
Despite the Assumptions & Simplifications, a robust agreement between simulation and experimental results is achieved at a Steady-State.

## Slide 32
/ Lung Deposition Pt.3
pMDI + Idealized Throat – Graphics
As particles impact to the wall – the tracking of such stops since the particles are considered trapped. Within Ansys Fluent software it is possible to locate the regions of particle impaction at the wall by using within Contours: Discrete Phase Variables… > DPM Wall Normal Pressure.
DPM Wall Normal Pressure captures the Impact intensity of Particles
And location of such. The graphics for 30 L/min is the following:

## Slide 33: Spacer Benefits

## Slide 34
The Spacer is a medical device that significantly improves API lung deposition by leveraging the rapid evaporation of volatile propellant components within its chamber. When propellant encounters rapid pressure drops and ambient temperature (25°C), it undergoes immediate flash evaporation, reducing droplet size significantly, since particle droplets consist of > 99% propellant & a remaining < 1% of non-volatile API.
Total Injected Mass: 0.0001 kg
Total Evaporated Mass: 0.00009 kg
/ Spacer Benefits Pt.1
Propellant Flashing & Droplet Size Reduction

## Slide 35
Particle Size Reduction and Flow Dynamics Literature demonstrates that droplets decrease substantially within the first 3 cm axially from spray release due to evaporation. By inhalation time, these smaller particles exhibit reduced relaxation time (proportional to particle diameter), enabling them to follow airflow patterns more readily with minimal gravitational settling. This reduces unwanted deposition in both the spacer chamber and throat.
/ Spacer Benefits Pt.2
Propellant Flashing & Droplet Size Reduction
~3 cm
Notes: Particle Filtration Mechanism Particles with incomplete evaporation (>5 μm) are effectively filtered by spacer deposition, preventing their delivery to airways where they would contribute to oropharyngeal rather than therapeutic lung deposition.

## Slide 36: pMDI + Spacer + Idealized Throat CAD Model

## Slide 37
/ pMDI + Spacer + Idealized Throat
CAD Model & Wall Named Selections
“wall_spacer”
“wall_mouth”
“wall_lower”
“wall_spacer” | “wall_mouth” | “wall_lower”
Plastic walls; Inhaler and Spacer Region | Oral Cavity, Pharynx and Larynx Regions | Trachea; Entry to lower respiratory system Region

## Slide 38: pMDI + Spacer + Idealized Throat CFD Model

## Slide 39
/ Particle Size
Using Species Transport, with initial Droplet Diameter of , an Injection Time of 0.1s and Mass Flow Rate of 0.0005 kg/s and desire for a final (Post-Evaporation) Droplet Diameter of , the following math is required:
Ansys Fluent software requires from the user: Initial Droplet Diameter ✔, and Mass Fraction for Each Component X
Using Cube-Law,
is non-volatile mass & is volatile mass
The total mass injected into the Domain is
And thus, the non-volatile mass is
NOTE: Common dosing within pMDI systems are 50-250 µg of API, but the non-volatile mass usually consist of an API and other solid residual, thus the API is just a fraction of the non-volatile component.
Mass Fraction for Each Component ✔
Components Mass Fractions | 
Volatile Component (Propellant) | Non-Volatile Component (API + Other)
0.992 | 0.008

## Slide 40
0
Check the Mesh Quality in: Domain > Quality > Evaluate Mesh Quality
Max. Aspect Ratio ~ 25
Minimum Orthogonal Quality ~0.25
/ Mesh
Quality Check

## Slide 41
/ Species Transport
Activation
To active the Species Transport Model go to Setup > Models > Species and Click on Species Transport & Click Edit… the Mixture Material
1
CFD
Create Fluid for Propellant as Vapor (Post Evaporation)
2
Notes: Research Objective: Optimize spacer geometry for maximum lung delivery
Computational Approach: Integrated CFD-DPM simulation
Design Variables: Spacer length and diameter
Success Metrics: Minimize wall deposition, maximize escaped particles

## Slide 42
/ Species Transport
Mixture
In Materials > Mixture & and edit the
Mixture Species
3
Add the “propellant-vapor” to selected Species and remove the default materials (“o2” & “h2o”)
4

## Slide 43
/ Cell Zone Conditions: Fluid
Changing material to Mixture
In Cell Zone Conditions > Fluid, make sure mixture is selected
5

## Slide 44
/ General
Gravity & Unsteady State/Transient
In Task Page > General > Time select Transient
6
In Task Page > Gravity (Check) and add Gravitational Acceleration as shown
7

## Slide 45
/ Discrete Phase Model (DPM)
Activation, Tracking, Physics Models & Numerics
In Models > DPM edit Particle Treatment, Tracking, Physical Models and Numerics as shown
8

## Slide 46
/ Discrete Phase Model (DPM)
Injection: Type & Point Properties
In Models > DPM > Injections create a new injection as shown in the RHS picture
9

## Slide 47
/ Discrete Phase Model (DPM)
Injection: Display

## Slide 48
/ Discrete Phase Model (DPM)
Injection: Physics Models & Turbulent Dispersion
In the injection created go to Physical Models and ensure it matches the picture below
10
Ensure Turbulent Dispersion matches
11

## Slide 49
/ Discrete Phase Model (DPM)
Injection: Components (Initial)
As DPM is initialized with Multicomponent, the Multicomponent Section will be available
Since no new particle materials have been created, the default option h2o is shown
Therefore, it is required to create new particle materials and add to the Particle Mixture so the mass fraction for API and Propellant can be defined within DPM
Initial
Desired

## Slide 50
/ Discrete Phase Model (DPM)
Droplet: Materials
In Materials > Droplet Particle and copy any 2 particle materials from Fluent software Database as they will be changed
12

## Slide 51
/ Discrete Phase Model (DPM)
Droplet: Materials – Propellant
Pick any of the to materials copied and change the Properties to create the Propellant Droplet Material as shown
13
Next Slide

## Slide 52
Temperature (K) | Temperature (°C) | Vapor Pressure (Pa)
200 | -73.15 | 1,013
220 | -53.15 | 10,130
240 | -33.15 | 51,330
247 | -26.15 | 78,900
260 | -13.15 | 167,000
280 | 6.85 | 371,000
300 | 26.85 | 770,000
320 | 46.85 | 1,470,000
/ Discrete Phase Model (DPM)
Droplet: Materials – Propellant; Saturation Vapor Pressure
Create 8 points in Saturation Vapor Pressure at given Temperatures with given Pressures from the table
14

## Slide 53
/ Discrete Phase Model (DPM)
Droplet: Materials – API
Pick the remaining material copied and change the Properties to create the API Droplet material as shown
15
Next Slide

## Slide 54
/ Discrete Phase Model (DPM)
Particle Mixture: Assign New Materials & Remove H2O
Like Mixture, In Particle Mixture add “propellant-liquid” and “api” to the selected Species and remove “water-liquid”
16

## Slide 55
/ Discrete Phase Model (DPM)
Injection: Components (Final)
Initial
Desired ✓
In DPM > Injection > Components the new particle materials/components will appear. Define the mass fractions:
“propellant-liquid”: 0.998
“api”: 0.002
Define Evaporating Species for the propellant, using “propellant-vapor” – created in Step 2
17

## Slide 56
/ Named Expressions
Inlet
Create a Named Expression for inhalation velocity (2.72 m/s) – inlet velocity and select as Input Parameter
18

## Slide 57
/ Boundary Conditions
Inlet: Momentum & Thermal
In Boundary Conditions > Inlet provide Velocity Magnitude as an Expression using the named expression “inlet_velocity”
19
In Thermal provide the Temperature as shown
20

## Slide 58
/ Boundary Conditions
Inlet: Species & DPM
In Species provide the Species Mass Fraction of “air” as shown
21
In DPM provide the Discrete Phase BC Type as escape
22

## Slide 59
/ Boundary Conditions
Outlet: Momentum & Thermal
In Boundary Conditions > Outlet ensure Gauge Pressure is 0
23
In Thermal provide the Temperature as shown
24

## Slide 60
/ Boundary Conditions
Outlet: Species & DPM
In Species provide the Species Mass Fraction of “air” as shown
25
In DPM provide the Discrete Phase BC Type as escape
26

## Slide 61
/ Boundary Conditions
Wall: Lower Tract; Thermal & DPM
In Thermal, select Temperature is Thermal Conditions and define the Temperature as shown
27
In DPM provide the Discrete Phase BC Type as trap
28

## Slide 62
/ Boundary Conditions
Wall: Upper Tract; Thermal & DPM
In Thermal, select Temperature is Thermal Conditions and define the Temperature as shown
29
In DPM provide the Discrete Phase BC Type as trap
30

## Slide 63
/ Boundary Conditions
Wall: Spacer Chamber; Thermal & DPM
In Thermal, select Temperature is Thermal Conditions and define the Temperature as shown
31
In DPM provide the Discrete Phase BC Type as trap
32

## Slide 64
/ Solution Methods
Pressure-Velocity Coupling & Spatial Discretization
Set the following Solution Methods in: Solution > Methods
33
Notes: Pressure-Velocity Coupling:
SIMPLEC – Faster pressure–velocity convergence in mildly transient DPM/species flows.
Spatial Discretization:
Gradient: Least Squares Cell-Based – Most accurate for unstructured meshes like your inhaler geometry.
Pressure: PRESTO! – Designed for strongly swirling or recirculating flows (ideal for inhaler jets/spacers).
Momentum: Second Order Upwind – Improves accuracy of velocity field without excessive diffusion.
Turbulent Kinetic Energy: Second Order Upwind – Captures turbulence gradients accurately in recirculating regions.
Specific Dissipation Rate: Second Order Upwind – Consistent with k-equation for smooth turbulence resolution.
Species (air, etc.): Second Order Upwind – Reduces numerical diffusion in species concentration fields.
Energy: Second Order Upwind – Ensures accurate thermal and enthalpy transport if temperature effects are active.

## Slide 65
/ Residual
Continuity, Momentum, Energy & Turbulence k-ꙍ SST
Set the following Residuals in: Solution > Monitors > Residuals
34

## Slide 66
/ Report Definitions
DPM: Injected Mass (injection)
Set the following DPM Report Definition in: Solution > Report Definitions > New > DPM Report
35

## Slide 67
/ Report Definitions
DPM: Evaporated Mass (injection)
Set the following DPM Report Definition in: Solution > Report Definitions > New > DPM Report
36

## Slide 68
/ Report Definitions
DPM: Escaped Mass (outlet)
Set the following DPM Report Definition in: Solution > Report Definitions > New > DPM Report
37

## Slide 69
/ Report Definitions
Summary
New Report Definitions

## Slide 70
/ Solution Initialization
Hybrid Initialization
Initialize the solutions in: Solution > Initialization > Hybrid Initialization
38

## Slide 71
/ Graphics
Plane Surface Creation
Create a new Plane Surface in the XY Plane in Graphics by Right-Clicking Surfaces > New > Plane…
38

## Slide 72
/ Graphics
Contour: Static Temperature
*Post Solution Image
Create a new Temperature Contour in Results > Graphics > Contours:
39

## Slide 73
/ Graphics
Particle Tracks: Particle Diameter
Create a new Particle Diameter Graphic, by Clicking Results > Graphics > Particle Tracks:
40
Track Style > Attributes…

## Slide 74
/ Graphics
Particle Tracks: Particle Temperature
*Post Solution Image
Create a new Particle Temperature Graphic, by Clicking Results > Graphics > Particle Tracks:
41

## Slide 75
/ Graphics
Mesh
Create a new Mesh Graphic, by Clicking Results > Graphics > Mesh:
42

## Slide 76
/ Graphics
Contour:
Create a new Contour Graphic, by Clicking Results > Graphics > Contours:
43

## Slide 77
/ Graphics
Contour: Turbulent Kinetic Energy
Create a new TKE Contour Graphic, by Clicking Results > Graphics > Contours:
44

## Slide 78
/ Graphics
Scene 1: Mesh + Particle Diameter
Create a new Particle Diameter + Mesh Scene, by Clicking Results > Scenes:
45
*Post Solution Image

## Slide 79
/ Graphics
Scene 2: Mesh + Particle Temperature
Create a new Particle Temperature + Mesh Scene, by Clicking Results > Scenes:
46
*Post Solution Image

## Slide 80
/ Graphics
Animation 1 & 2 from Scene 1 & 2
47
Create a new Solution Animations for both Scenes in Solution > Calculation Activities:

## Slide 81
/ Solution
Run Calculations
NOTE: This simulation is computationally expensive and may take hours to days depending on time-step, Number of Time Steps and System capabilities.
To reduce run time on modest systems, decrease the number of time steps or increase the time-step size and reduce output frequency — but expect reduced temporal resolution and less-accurate particle trajectories/animations.
For Reference

## Slide 82: Results & Post-Processing CFD Model

## Slide 83
/ Particle Diameter
Animation
Check the Folder for the particle-diameter-animation.mp4!

## Slide 84
/ Particle Temperature
Animation
Check the Folder for the particle-temperature-animation.mp4!

## Slide 85
/ Particle Tracking
Summary
By Selecting Summary in Particle Tracks < Graphics, The following information will be displayed in the Console

## Slide 86
/ Particle Tracking
Summary
After a period of 1 second – 10000 Number of Time Steps of Size 0.0001 [s], the Summary is the following:
Results @ t = 1s |  |  |  |  |  | 
Flow Rates () | Injected Particles | Trapped Particles | Particle In Transit | Escaped Particles | Model Efficiency | Experimental Efficiency
30 | 50000 | 7563 | 6637 | 35800 | 71% | 70-80%
Results @ t = 1s |  |  |  |  |  | 
 | Injected Particles | Trapped Particles | Particle In Transit | Escaped Particles | Model Efficiency | Experimental Efficiency
30 | 50000 | 7563 | 6637 | 35800 | 71% | 70-80%
The simulation results demonstrate a substantial improvement in aerosol delivery performance with the use of a spacer device. Lung deposition efficiency increased to 71%, corresponding to roughly a 20% enhancement relative to the pressurized metered-dose inhaler (pMDI) coupled directly to the idealized throat model. This improvement aligns with expected trends, as the spacer reduces particle inertia and impaction losses in the oropharyngeal region, allowing finer droplets to enter the lower airways.

## Slide 87
/ Total Escaped Mass vs Flow Time
Summary
Even with a 1-second run, the simulation reached 71% lung deposition, with some particles still in transit.
This shows great progress — extending the simulation time would only improve result accuracy and complete particle tracking:
Question: What do you expect to happen to the total number of escaped particles if the simulation continues?
Hint: The curve should eventually plateau, as remaining particles either deposit on surfaces or exit the domain.

## Slide 88
/ Conclusions
This simplified step-by-step tutorial establishes a strong foundation for pMDI drug delivery modeling in Ansys Fluent software, incorporating key physical processes such as transient flow, species transport, droplet evaporation, and discrete phase tracking.
Building on this framework, users can progressively extend the model to include additional physical phenomena for enhanced realism and predictive accuracy, such as:
Droplet Break-up Modeling – Capturing secondary atomization and surface tension effects based on formulation properties.
Advanced Deposition Models – Implementing probabilistic or velocity-dependent sticking efficiencies for more accurate wall interactions.
Electrostatic Effects – Accounting for particle charging and image forces, especially relevant for plastic spacer geometries.
Hygroscopic Growth – Modeling particle size variation due to humidity uptake under respiratory conditions.
Together, these extensions enable a more comprehensive simulation of aerosol behavior, supporting both research-level investigations and device optimization studies in pulmonary drug delivery.

## Slide 89
Kleinstreuer C, Zhang Z. Computational analyses of a pressurized metered dose inhaler and a new drug-aerosol targeting methodology. J Aerosol Med. 2003;16(3):249-261. Available from
Golshahi L, Finlay WH. An idealized child throat that mimics average pediatric oropharyngeal deposition. Aerosol Sci Technol. 2012;46(5):i–iv. doi:10.1080/02786826.2012.667170 Available from
Vincken W, Levy ML, Scullion J, Usmani OS, Dekhuijzen PNR, Corrigan CJ. Spacer devices for inhaled therapy: why use them, and how? ERJ Open Res. 2018 Jun 18;4(2):00065-2018. doi: 10.1183/23120541.00065-2018. Available from
Cheng YS. Mechanisms of pharmaceutical aerosol deposition in the respiratory tract. AAPS PharmSciTech. 2014 Jun;15(3):630-40. doi: 10.1208/s12249-014-0092-0. Epub 2014 Feb 22. PMID: 24563174; PMCID: PMC4037474. Available from
F. Huang, Y. Zhang, Z.B. Tong, X.L. Chen, R.Y. Yang, A.B. Yu, Numerical investigation of deposition mechanism in three mouth–throat models, Powder Technology, Volume 378, Part A, 2021, ISSN 0032-5910, https://doi.org/10.1016/j.powtec.2018.11.095.
/ References

## Slide 90: Ansys Education Resources Feedback Survey
Here at Ansys, we rely on your feedback to ensure the educational content we create is up-to-date and fits your teaching needs.
Please click the link below to fill out a short survey (~7 minutes) to help us continue to support academics around the world utilizing Ansys tools in the classroom.
Feedback Survey Link
Notes: Feedback Survey link for copy-paste method  https://ansys.typeform.com/to/jcattJI5